## Sheet1
Last updated 9/2020
Manufacturer | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Stirling Ultracold | Panasonic | Panasonic | Panasonic | Panasonic
Freezer Model | TSX60086A | TSX70086A | TSX50086A | TSX40086A | SU780XLE | MDF-DU702VH-PA | MDF-DU702VHA-PA | MDF-DU502VH-PA | MDF-DU502VHA-PA
Capacity (L) | 816 | 949 | 682 | 549 | 780 | 729 L | 729 L | 528 L | 528 L
Volume (cu.ft). | 28.8 | 33.5 | 24.1 | 19.4 | 27.5 | 25.7 | 25.7 | 18.6 | 18.6
2in. Box Capacity | 600 | 700 | 500 | 400 | 600 | 576 | 576 | 384 | 384
Exterior (H x D x W) | 78 x 37.8 x 39 in | 78 x 37.8 x 45.2 in | 78 x 37.8 x 34 in | 78 x 37.8 x 28.3 in | 78.5 x 34.3 x 36 in. | 77.25 x 33.75 x 40.62 | 78.5 x 36.8 x 40.6 | 78.5 x 34.7 x 31.1 | 78.5 x 34.7 x 31.1
Exterior (H x D x W) | 1981 x 960 x 1006 mm | 1981 x 960 x 1148 mm | 1981 x 960 x 864 mm | 1981 x 960 x 719 mm | 1994 x 870 x 915 mm | 1962 x 857 x 1032 | 1993 x 935 x 1030 | 1993 x 882 x 790 | 1993 x 882 x 790
Floor Footprint (m2) | 0.9 | 1.1 | 0.83 | 0.69 | 0.79605 | 0.88442 | 0.91 | 0.7 | 0.7
Average kWh | 9.4 | 10.4 | 8.7 | 8.3 | 7.6 | 8 | 7.453 | 6.51 | 6.882
Energy Star Listing | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes
Steady State Energy Use (ENERGY STAR® Final Test Method) | 7.4 kW-hr/day | 10.4 kW-hr/day | 8.7 kW-hr/day | 8.3 kW-hr/day | 6.86 kWh/day at -75° C (Weighted Average) | 7.295 kWh/day at -75° C (Weighted Average) | No Data | 5.975 kWh/day at 75° C (Weighted Average) | No Data
Average kWh/day/cu-ft | 0.3264 | 0.36 | 0.45 | 0.49 | 0.2495 | 0.31 | 0.29 | 0.35 | 0.37
Temperature Range | -80C to -50C adjustable to 1C increments | -80C to -50C adjustable to 1C increments | -80C to -50C adjustable to 1C increments | -80C to -50C adjustable to 1C increments | -86C to -20C, adjustable to 1C increments | -86C to -50C adjustable to 1C increments | -86 to -50C adjustable to 1C increments | -86 to -50C adjustable to 1C increments | -86 to -50C adjustable to 1C increments
Electric Power | 115V, 60 Hz, 1 Phase | 115V, 60 Hz, 1 Phase | 115V, 60 Hz, 1 Phase | 115V, 60 Hz, 1 Phase | 100V to 240V (±10%) at either 50 or 60Hz | 208-230V | 115V | 220V | 115V
Electrical Plug | NEMA 5-15P Plug / IE Cord 10 ft (3.05m) | NEMA 5-15P Plug / IE Cord 10 ft (3.05m) | NEMA 5-15P Plug / IE Cord 10 ft (3.05m) | NEMA 5-15P Plug / IE Cord 10 ft (3.05m) | NEMA 5-15P (120V) or NEMA 6-15P (240V) | NEMA 6-15P (208/230 V) | NEMA 5-20P (115V) | NEMA 6-15P (220V) | NEMA 5-20P (115V)
Certification/Agency Listing | UL, cUL | UL, cUL | UL, cUL | UL, cUL | UL, cUL, CE | N/A | N/A | N/A | N/A
Security | Pad Lock, Optional Acess Card Entry | Pad Lock, Optional Acess Card Entry | Pad Lock, Optional Acess Card Entry | Pad Lock, Optional Acess Card Entry | Door Lock, Optional PIN requirement built in | Key Lock, optional pad lock | Key Lock, optional pad lock | Key Lock, optional pad lock | Key Lock, optional pad lock
Event Log | DevieLink Come Standard Datalogging & Temperature Mapping, Export via USB Port | DevieLink Come Standard Datalogging & Temperature Mapping, Export via USB Port | DevieLink Come Standard Datalogging & Temperature Mapping, Export via USB Port | DevieLink Come Standard Datalogging & Temperature Mapping, Export via USB Port | All alarms, door openings | All alarms, door openings | All alarms, door openings | All alarms, door openings | All alarms, door openings
Refrigeration System | Two Stage Cascade System - Variable Speed Compressor | Two Stage Cascade System - Variable Speed Compressor | Two Stage Cascade System - Variable Speed Compressor | Two Stage Cascade System - Variable Speed Compressor | Helium charged free-piston Stirling engine with continuous modulation, Gravity driven thermosiphon | Dual Compressor Cascade System | Dual Compressor Cascade System | Dual Compressor Cascade System | Dual Compressor Cascade System
Insulation | Vacuum Insulation Panels with High-Density Water-Blown Polyurethane Foam | Vacuum Insulation Panels with High-Density Water-Blown Polyurethane Foam | Vacuum Insulation Panels with High-Density Water-Blown Polyurethane Foam | Vacuum Insulation Panels with High-Density Water-Blown Polyurethane Foam | High performance vacuum insulated panels and non-HFC polyurethane foam using the Ecomate® environmentally friendly blowing agent | Vacuum insulated panels | Vacuum insulated panels | Vacuum insulated panels | Vacuum insulated panels
Noise dB (A) | 48 | 48 | 48 | 47.4 | 45 | 46 | 46 | 46 | 46
Shelves | 3 adjustable stainless steel - Configurable shelves to fit your storage need | 3 adjustable stainless steel - Configurable shelves to fit your storage need | 3 adjustable stainless steel - Configurable shelves to fit your storage need | 3 adjustable stainless steel - Configurable shelves to fit your storage need | 2 adjustable stainless steel (standard), additional shelves optional | 3 | 3 | 3 | 3
Inner Doors | 2 insulated stainless steel with latches | 2 insulated stainless steel with latches | 2 insulated stainless steel with latches | 2 insulated stainless steel with latches | 3 insulated with magnetic latches | 2 (insulated; ABS, with stainless steel frames and positive latches) | 2 (insulated; ABS, with stainless steel frames and positive latches) | 2 (insulated; ABS, with stainless steel frames and positive latches) | 2 (insulated; ABS, with stainless steel frames and positive latches)
Optional Accessories | DeviceLink Temp. Monitoring come standard, LN2 & CO2 backup, chart recorder | DeviceLink Temp. Monitoring come standard, LN2 & CO2 backup, chart recorder | DeviceLink Temp. Monitoring come standard, LN2 & CO2 backup, chart recorder | DeviceLink Temp. Monitoring come standard, LN2 & CO2 backup, chart recorder | Freezer Racks (sold separately), Chart recorder, CO2 and LN2 back-up systems, additional shelves, 4-20mA temperature output | LabAlert® Monitoring System, Freezer Racks (sold separately), Chart recorder, CO2 and LN2 back-up systems | LabAlert® Monitoring System, Freezer Racks (sold separately), Chart recorder, CO2 and LN2 back-up systems | LabAlert® Monitoring System, Freezer Racks (sold separately), Chart recorder, CO2 and LN2 back-up systems | LabAlert® Monitoring System, Freezer Racks (sold separately), Chart recorder, CO2 and LN2 back-up systems
Warranty | 5 year parts & labor, Additional 7 year on Variable Speed Compressors | 5 year parts & labor, Additional 7 year on Variable Speed Compressors | 5 year parts & labor, Additional 7 year on Variable Speed Compressors | 5 year parts & labor, Additional 7 year on Variable Speed Compressors | 2 years Parts and Labor (7-years Engine and Thermosiphon Protection) | 5 year parts and labor | 5 year parts and labor | 5 year parts and labor | 5 year parts and labor
Data Source | https://www.thermofisher.com/order/catalog/product/TSX60086A#/TSX60086A | https://www.thermofisher.com/order/catalog/product/TSX60086A#/TSX60086A | https://www.thermofisher.com/order/catalog/product/TSX60086A#/TSX60086A | https://www.thermofisher.com/order/catalog/product/TSX60086A#/TSX60086A | https://www.stirlingultracold.com/wp-content/uploads/2018/03/20010_GLOBAL_SU780XLE_Technical_Data_Sheet_Web.pdf | https://www.phchd.com/us/biomedical/preservation/ultra-low-freezers/mdf-du702vhpa | https://www.phchd.com/us/biomedical/preservation/ultra-low-freezers/mdf-du702vhapa | https://www.phchd.com/us/biomedical/preservation/ultra-low-freezers/mdf-du702vhapa | https://www.phchd.com/us/biomedical/preservation/ultra-low-freezers/mdf-du702vhapa